[image: image1.png]WOODLEY

EQUIPMENT COMPANY
ﬁ




MAINTENANCE PROTOCOL – 
WOODLEY abc HAEMATOLOGY ANALYSER

DAILY (See Section 9 abc Manual)

1.Disconnect Analyser from mains supply.

   Wipe off any blood spills from the outside of the analyser with warm water.

   NB. NEVER USE SOLVENT OR ABRASIVE MATERIALS

2.Open door and inspect interior of the analyser.

   Wipe off any blood / reagent splashes with warm water.

   Check for any leaks or crystalline build-ups – report as necessary.

3.Check that the pack is not distended with air build-up – “burp” if necessary.

4.Close door.

   Switch on the analyser.

   Check that “Start-up” is passed.

WEEKLY

Run “Minotrol” Control Fluid using “Control” Species Card.

Check all parameters are within tolerance limits.

Adjust Calibration Coefficients as necessary (See Section 7 ABC Manual).

EVERY 100 SAMPLES (Or monthly according to usage)

Run “Concentrated Clean” (See Section 9 abc Manual).

EVERY 6 MONTHS

Run “Autocalibration” using Minotrol (See Section 7 abc Manual).

NB.REPLACING REAGENT PACKS

When installing a new Reagent Pack it is advisable to Prime the analyser with the Waste Pipe feeding into a deep open container (eg. a 500ml beaker) – this will prevent a lot of air entering the waste chamber of the pack.

When Priming is complete connect the waste pipe to the pack.

IF IN DOUBT CONSULT THE ABC MANUAL

NB. NEVER RUN CLOTTED SAMPLES THROUGH THE ABC – THIS WILL BLOCK THE ANALYSER AND IS NOT COVERED BY THE WARRANTY OR YOUR SERVICE AGREEMENT.

